## Sheet2
Equipment Name | Approved Abbreviation | Location | Make (Manufacturer) | Model | Serial No | Funding Agency Type | Funding Agency Details (Use Comma if multiple) | Equipment Description | Department/Lab Name | Project Title | Rate/Hr | Refernce Website | Category
Powder X-Ray Diffraction unit | Powder XRD | ARCI-Chennai | Rigaku-Japan | Smart Lab | AR2001 | Internal | ARCI | The high-resolution X-ray diffractometer provided with SmartLab Guidance software, acts as an intelligent interface for the user. The system incorporates a high resolution ?/? closed loop goniometer drive system, cross beam optics (CBO), an in-plane scattering arm, and an optional 3.0 kW rotating anode generator. High flexibility by coupling a computer controlled alignment system with a fully automated optical system makes it easy to switch between hardware modes. Optional in-plane diffraction arm for in-plane measurements without reconfiguration and focusing and parallel beam geometries without reconfiguration. | CENTRE FOR FUEL CELL TECHNOLOGY (CFCT) |  |  |  | Characterization
Atomic Force Microscope | AFM | ARCI-Hyderabad | Park systems-South Korea | Park XE7 (Direct On-Axis Manual Focus Optics) | AR2002 | Internal | ARCI | The atomic force microscope (AFM) is extensively used in material science and biological sciences. Characterization of sol-gel thin film surface morphology in sub-micron scale is necessary to evaluate their function properties in a wide variety of applications. It has been known that sub-micron and nanoscale properties control various aspects of functional performance. For example, enhanced transmittance in borosilicate glass (BSG, 91%) over soda lime glass (SLG, 89%) is due to sub-nanometer surface morphology of BSG. This may occur due to its different chemical composition though the refractive indices are comparable. | CENTRE FOR SOL-GEL COATINGS (C-SOL) |  |  |  | Characterization
Carbon-Sulphur Analyzer | C-S Analyzer | ARCI-Hyderabad | Leco-USA | CS - 444, | AR2003 | Internal | ARCI | Carbon/Sulphur system is a microprocessor based software driven instrument for wide range measurement of Carbon and Sulphur content of metals, ores, ceramics and other inorganic materials. It uses HF-400 induction furnace (18 MHz, 2,2 kW) for CO2/SO2 conversion. It works on the principle of infrared absorption of CO2 and SO2. | CENTRE FOR NANOMATERIALS (CNM) |  |  |  | Characterization
Cyclic corrosion Chamber |  | ARCI-Hyderabad | Atlas-USA | Atlas-USA | AR2004 | External | General Motors, USA | To better understand the potential for adverse effects of outdoor exposure, materials are regularly tested to various environmental conditions. The cyclic corrosion cabinet is now accepted as superior at duplicating actual environments. It improves repeatability, reduces time spent in handling samples and also operator error by creating different environments in one cabinet. Standard exposure cycles like salt or chemical fog, water fog, salt spray, dry-off cycle, high temperatures i.e. up to 70oC can be increased for required time from hours to number of days which are closer to environmental changes. | CENTRE FOR MATERIALS CHARACTERIZATION AND TESTING (CMCT) |  |  |  | Characterization
Electrochemical Corrosion Testing Equipment |  | ARCI-Hyderabad | Soartron-UK | Soartron-UK | AR2005 | Internal | ARCI | The corrosion is an electrochemical process and electrochemical potential is the driving force for reactions. The principle employed in the electrochemical interface is based on the application of a controlled voltage to a sample (working electrode) and measurement of the change in current, or vice versa. The corrosion rates of the specimens can be determined using CorrWare programme. Most of the corrosion types uniform, localized, galvanic, dealloying; stress corrosion and hydrogen-induced failure can be analyzed by this technique. Conventional electrochemical experiments like static, dynamic polarization, potentiostatic, potentiodynamic experiments etc. can be done. This system is extensively used to study the corrosion properties of coated substrates. Based on the results, further improvement of coatings has been done. Different types of coatings like DLC on Al alloys, MAO coatings on Al, Coating on steels and Al alloys by Sol-Gel technique, materials like Mg-Zn, steels, alloys have been analyzed. The system can also be used to study batteries, fuel cells, inhibitors, electroplating and catalytic properties etc. | CENTRE FOR MATERIALS CHARACTERIZATION AND TESTING (CMCT) |  |  |  | Characterization
Dual beam Microscope |  | ARCI-Hyderabad | Zeiss GMBH-Germany | Zeiss GMBH-Germany | AR2006 | Internal | ARCI | This system is equipped with an electron beam and a focused ion beam (FIB). The Ga-ion beam is used to remove material from the sample surface in a controlled fashion and the electron beam is used to carry out the imaging. As the FIB unit is present in-situ in the FE-SEM, live imaging is possible. Applications include study of coating growth mechanisms, corrosion etc. The system can also be used to prepare samples for EBSD, where a mirror finish and relatively strain-free surface is necessary. There are several ports available for additional equipment to be installed such as a nanomanipulator that can be used as a TEM sample preparation tool, a nanolithography system that can be used for preparation of samples for microtensile testing and so on. If an EBSD unit is attached to the dual beam microscope, the three-dimensional extent of grains in the sample can be mapped out. | CENTRE FOR MATERIALS CHARACTERIZATION AND TESTING (CMCT) |  |  |  | Characterization
Erosion Wear Test Rig |  | ARCI-Hyderabad | ARCI-India | ARCI-India | AR2007 | Internal | ARCI |  | CENTRE FOR ENGINEERED COATINGS (CEC) |  |  |  | Characterization
FTIR Spectrophotometer For Spectral And Thermal Emissivity Measurements |  | ARCI-Hyderabad | Bruker Optik-Germany | Bruker Optik-Germany | AR2008 | Internal | ARCI |  | CENTRE FOR SOLAR ENERGY MATERIALS (CSEM) |  |  |  | Characterization
Gas Chromatograph |  | ARCI-Chennai | Thermo FIsher Scientofic-USA | Thermo FIsher Scientofic-USA | AR2009 | Internal | ARCI | Gas chromatograph with TCD, FID and FPD detectors is used to analyse the exhaust gas from the fuel cell especially, when the fuel cell is operated with methanol and reformate gas. | CENTRE FOR FUEL CELL TECHNOLOGY (CFCT) |  |  |  | Characterization
Instrumented Indentation Testing Facility |  | ARCI-Hyderabad | Micro Materials-UK | Micro Materials-UK | AR2010 | Internal | ARCI |  | CENTRE FOR ENGINEERED COATINGS |  |  |  | Characterization
Micro Raman Spectrometry |  | ARCI-Hyderabad | Jobin Yuon-France | Jobin Yuon-France | AR2011 | Internal | ARCI | Raman Spectrometery is complementary to FTIR. Raman can be used to analyse aqueous solutions since it does not suffer from the large water absorption effects found with FT techniques. Raman requires little or no sample preparation. Raman Spectroscopy actually originates from the bond vibrations. This intrinsic nano probing combined with micro raman spectroscopy makes it very sensitive to short range structure in the nano materials. Thus it is good complement to TEM or X-Ray Diffraction techniques. The advantage of No sample preparation required both for solid and liquid samples in Micro Raman Spectroscopy makes it an easy access to characterize nano materials. Micro Raman provides phase identification, even traces of secondary phase can be detected as the polarisibility changes for different kinds of bonds which is detected in the Raman intensity. | CENTRE FOR NANOMATERIALS (CNM) |  |  |  | Characterization
Microhardness Tester |  | ARCI-Hyderabad | Walter UHL-Germany | Walter UHL-Germany | AR2012 | Internal | ARCI | With the aid of computer control, arrays of indents can be made on the sample and the hardness can be measured as a function of distance from a reference point. This unit is especially suitable to induce cracks on the sample surface so that the fracture toughness can be measured. | CENTRE FOR MATERIALS CHARACTERIZATION AND TESTING (CMCT) |  |  |  | Characterization
Nanosizer |  | ARCI-Hyderabad | Malvern-UK | Malvern-UK | AR2013 | Internal | ARCI | The instrument measures the diffusion coefficient (D) and converts this to particle size, using the Stokes-Einstein equation: D=kT/3S, D=Hydrodynamic diameter, T=Absolute temperature, D=Diffusion Coefficient, K=Boltzman's constant, =Viscosity/c | CENTRE FOR CERAMIC PROCESSING (CCP) |  |  |  | Characterization
Non-Contact Optical Profilometer |  | ARCI-Hyderabad | Zygo-USA | Zygo-USA | AR2014 | Internal | ARCI | The system is used for estimating surface roughness and other surface parameters, as also to create 3D profiles of surfaces. It uses white light interferometry to map specimen surfaces. A beam splitter is used to split the incoming beam into two. One part goes to a reference mirror that is optically smooth while the other part scans the sample surface. The interference pattern between the two beams (that is, the one coming from the reference mirror and the other from the sample) is used to recreate the surface profile. The resolution of the unit is 0.1 nm and hence optically smooth surfaces can be studied. An advantage of this system over an atomic force microscope is that there is no mechanical contact with any part of the sample and hence even smooth and soft surfaces can be profiled. Another advantage is that the optics stay much above the sample and hence even deeper areas of the sample surface can be imaged as far as the light beam is able to fall on the area of interest. | CENTRE FOR MATERIALS CHARACTERIZATION AND TESTING (CMCT) |  |  |  | Characterization
Particle Size Analyzer |  | ARCI-Hyderabad | Malvern-UK | Malvern-UK | AR2015 | Internal | ARCI | TRange: 0.02 micron-2000 micron | CENTRE FOR CERAMIC PROCESSING (CCP) |  |  |  | Characterization
Residual Stress Analyzer |  | ARCI-Hyderabad | Panalytical-Netharlands | Panalytical-Netharlands | AR2016 | Internal | ARCI | Capability: dry and wet | CENTRE FOR ENGINEERED COATINGS (CEC) |  |  |  | Characterization
Rheometer |  | ARCI-Chennai | Brookfield-USA | Brookfield-USA | AR2017 | Internal | ARCI | Solvent condition: aqueous and non-aqueous,Measures particle size distribution and size | CENTRE FOR FUEL CELL TECHNOLOGY (CFCT) |  |  |  | Characterization
Scanning Electron Microscope (SU1510) With EDS |  | ARCI-Chennai | Hitachi-Japan | Hitachi-Japan | AR2018 | Internal | ARCI |  | CENTRE FOR FUEL CELL TECHNOLOGY (CFCT) |  |  |  | Characterization
Scanning Electron Microscope (S3400) With EDS |  | ARCI-Hyderabad | Hitachi-Japan | Hitachi-Japan | AR2019 | Internal | ARCI |  | CENTRE FOR MATERIALS CHARACTERIZATION AND TESTING (CMCT) |  |  |  | Characterization
Scanning Electron Microscope (Gemini 500) With EDS/EBSD |  | ARCI-Hyderabad | Zeiss GMBH-Germany | Zeiss GMBH-Germany | AR2020 | Internal | ARCI | his instrument is used for studying sample surfaces at high magnifications (> 200000x) due to the presence of a hot Schottky field emission (FE) gun. Potential (called the extraction voltage) is applied to draw out the electrons. Due to this combination of high temperature and extraction voltage, the brilliance of the source beam, as also its stability, is high. This system is attached with an EDS unit and an electron back scatter diffraction (EBSD) unit. | CENTRE FOR MATERIALS CHARACTERIZATION AND TESTING (CMCT) |  |  |  | Characterization