## Sheet2
Equipment Name | Approved Abbreviation | Location | Make (Manufacturer) | Model | Serial No | Funding Agency Type | Funding Agency Details (Use Comma if multiple) | Equipment Description | Department/Lab Name | Project Title | Rate/Hr | Refernce Website | Category
Simultaneous Thermal Analyser |  | ARCI-Hyderabad | NETZSCH-Germany | NETZSCH-Germany | AR2021 | Internal | ARCI |  | maths |  |  |  | Characterization
Surface Area And Porosity Analyser |  | ARCI-Hyderabad | Micromeritics-USA | Micromeritics-USA | AR2022 | Internal | ARCI | Surface area and porosity are two important physical properties of the nanopowders that can be measured by this method. The equipment works on gas absorption principle therefore, it is very effective tool to study the surface morphology of the materials. Surface area (as low as 0.05m2/g) can be measured. It can measure adsorption and desorption isotherms to determine the shape of pores. It is possible to carryout measurements using adsorbates such as N2, Ar, CO2and Kr selectively and also capable of measuring Langmuir surface area and pore size distribution. Volume, area, total pore volume and pore size in the range 0.35 to 500nm can be determined. | Physics |  |  |  | Characterization
Thermogravimetry Analyser With Mass Spectrometer |  | ARCI-Chennai | NETZSCH-Germany | NETZSCH-Germany | AR2023 | Internal | ARCI | It consists of a vacuum-tight thermo micro balance system, microprocessor-controlled, with integrated calibration weight, with water-cooled heater, exchangeable sample carrier including radiation shield, programmable gas flow for 3 gases and integrated heating power supply, automated lid lifting device (prepared autovac for automated vacuum generation/ventilation). This instrument comes with QMS coupling for evolved gas analysis from 1-300 amu and a processing module of the Aeolos software (as of version 7.02), enabling a triggered start and stop as well as the transfer of the temperature signal into the Aeolos software for programming and control. unit for data acquisition and processor of max. 64 measuring channels (mass and mass range resp.) | Biology |  |  |  | Characterization
Thin Film Analyser |  | ARCI-Hyderabad | Filmetrics-USA | Filmetrics-USA | AR2024 | Internal | ARCI | The thin film analyzer works based on the intensity measurement of reflecting and transmitting light through the sample. Hence, thickness of only transparent coatings can be measured, though the substrates can be of transparent or non-transparent nature. There are possibilities of using a contact or a non-contact probe depending upon the nature of film and the substrate. | Biology |  |  |  | Characterization
Transmission Electron Microscope |  | ARCI-Hyderabad | FEI-Netherlands | FEI-Netherlands | AR2025 | Internal | ARCI | This instrument is used for studying samples at very high magnifications (> 500000X) in transmission. Samples of 3mm diameter are prepared such that the thickness at the centre is 200 nm or less so that electron transparency can be achieved. The incident electron beam can be accelerated to 200 kV and on passing through the sample, several interactions take place. The transmitted beam is then used for imaging and also electron diffraction studies. The system is equipped with an LaB6 filament providing a high beam current and long filament life. The unit has EDS and Electron Energy Loss Spectroscopy attachments, both of which are techniques for elemental analysis. In the later technique, the energy lost by the incident beam on passing through the sample is measured and thus elemental information can be obtained. | Biology |  |  |  | Characterization
UV-Vis NIR Spectrometer |  | ARCI-Hyderabad | Shimandzu-Japan | Shimandzu-Japan | AR2026 | Internal | ARCI |  | Biology |  |  |  | Characterization
UV Visible Spectrometer |  | ARCI-Hyderabad | LAMBDA | LAMBDA | AR2027 | Internal | ARCI |  | Biology |  |  |  | Characterization
Small Angle Xray Scattering (SAXS) |  | ARCI-Hyderabad | XENOCS-France | XENOCS-France | AR2028 | Internal | ARCI | A laboratory small-angle X-ray scattering (SAXS) system with a dual energy (Mo and Cr) source has been designed with the aim of studying the heterogeneity in high-Z metallic structural materials such as steels (Fe alloy) and other materials. With the combination of Mo and Cr energies, various camera lengths (maximum of 2400 mm), and an area detector, three decades of q range has been achieved, from 0.024 to 14 nm-1. In real space, the probing periodic distance is a maximum of 261 nm. In addition to having Mo and Cr sources, a flexible sample mounting stage permits measurements on a wide range of materials in transmission geometry. More details about instrument can be found at “A multi-functional dual-energy laboratory Mo-Cr-SAXS system” J. Appl. Cryst. (2015). 48, 2040-2043 https://doi.org/10.1107/S1600576715018804 http://scripts.iucr.org/cgi-bin/paper?S1600576715018804 | Biology |  |  |  | Characterization
X-Ray Photoelectron Spectrometer |  | ARCI-Hyderabad | OMICRON-Germany | OMICRON-Germany | AR2029 | Internal | ARCI | The instrument has the capabilities of depth profiling by Ar ion etching, Imaging XPS, Auger photoelectron spectroscopy (static and scanning modes), Scanning Electron Microscopy (SEM), Sample heating (upto 100oC) and cooling (liq N2 temp.). The technique being surface sensitive, it is very useful to characterize nanomaterials including, powders and thin films. | Biology |  |  |  | Characterization
Universal Testing Machine |  | ARCI-Hyderabad | INSTRON-UK | INSTRON-UK | AR2030 | Internal | ARCI |  | Biology |  |  |  | Characterization
Micro Focus X-Ray Diffractometer |  | ARCI-Hyderabad | RIGAKU-Japan | RIGAKU-Japan | AR2031 | Internal | ARCI | This is the most versatile laboratory scale x-ray diffraction system. It is equipped with high intensity microfocus rotating anode x-ray source-Rigaku MicroMax 007HF, which has brightness close to a second generation synchrotron with x-ray flux of 1014 to 1015 x-ray photon/mm2/s at the focal point. The system also has a highly sensitive and large image plate based 2-dimensional (2D) detector, which ensures a full scan from -47 to +163 degree for 2theta in a single exposure of very small duration. It has both Cu and Cr target option with provision to bring the beam spot size down to 10 micron. The system can be operated in reflection, transmission and glancing incidence configuration. It can perform wide range of studies like phase analysis, texture, residual stress, trace phase detection (for phase fraction of 0.1% or less) in micro/macro-area apart from thin film analysis in the glancing incidence mode. In addition, it has a provision to perform fully automated area mapping using auto stage. | Biology |  |  |  | Characterization
Scanning Electron Microscope (Merlin Compact) With EDS |  | ARCI-Chennai | Zeiss GMBH-Germany | Zeiss GMBH-Germany | AR2032 | Internal | ARCI | High resolution Field Emission Gun scanning electron microscope (FEG-SEM), (Model: Merlin Compact, Make: Zeiss, Germany) is available at ARCI, Chennai. The instrument has a best resolution of 0.8 nm with accelerating voltage variable from 0.02 to 30 keV. The FEG source is of thermionic emission type. The SEM is equipped with Secondary electron detector, In-lens detector, Back scattered Electron detector and EDS detector (EDAX, USA) for compositional analysis. The sample stage is a 5 axis euccentric stage with a tilt angle ranging from -3 to 70 0. The Gemini I column enables high resolution and high probe current making it more suitable for analytical applications. | maths |  |  |  | Characterization