## Sheet1
Specifications
Item : 12 in 1 Multiple Test and measuring system
1. Scope of Supply : Supply of software configurable hardware platform for twelve test and measuring instrument FPGA based with additional functionalities as per the detailed technical specifications.
2. Detailed Technical Specifications:
| Description / Requirement | Value / Remark
| Number of Inputs | ≥ 2 Nos
| Input Bandwidth | Shall cover the range DC to 200 MHZ
| Input Sampling Rate | ≥ 500 MS/s
| Input Resolution | ≥ 12 bits
| Input Impedance | Selectable 50Ohms / 1MOhms
| Input Coupling | Selectable AC / DC
| Input referred Noise (>1MHz) | ≤ 30 nV/√Hz
| Maximum Input Voltage | 10Vpk-pk
| Number of Outputs | ≥ 2Nos
| Output Bandwidth | Shall cover the range DC to 300 MHz
| Output Sampling Rate | ≥ 1GSamples/S
| Output Resolution | ≥ 16 bits
| Output Impedance | 50 Ohms
| Reference Clock | Selectable internal / external
| Reference Clock Frequency | 10 MHz
| Reference Clock Stability | ≤ 500 ppb
| Wireless Display | ≥ 10 inch display with multi touch for wireless configuration and monitoring
| Storage | : to SD ,RAM or cloud
| Software Interfaces (API) | Shall provide Application Programming Interface for communicating with device using MATLAB, Python, Labview.
1 | Oscilloscope Feature
| No of channels | 2
| Bandwidth | 200 MHz
| Sampling rate | 500 MS/s
| Input range | ,+/- 5 V
| Integrated Waveform Generator | 250 MHz
| Waveforms | sine,Square,ramp,pulse,DC
| Math channel | Add,subtract,multiply,divid.XY Mode,FFT,Arbitraray equation
| Coupling | AC/DC
| Input impedance | 50/1 M ohms
| Visualisation tools | Graphical/Numerical/Trend and histogram
2 | Data logger features include
| No of channels | 2
| Acquisition rate | upt to 1 MS/sec
| Input range | ,+/- 5V
| Max sampling rate into RAM | 1M S/sec for 1 channel or 500KS/s for two channels
| Max sampling rate into SD | 100KS/sec
| Storage | Dropbox,Email,icloud
| Format | CSV or binary format
3 | PID Controller Function Features | Shall have at least the following features
1. Two independent PID controllers.
2. User selectable gain profile: Proportional (P), Integral (I), Differential (D), Double Integral
3. Settable Proportional Gain shall cover the range -60 dB to 60 dB
4. Settable Integrator Crossover frequency shall cover the range 1Hz to 100 kHz
5. Settable Differentiator Crossover frequency shall cover the range 10 Hz to 1MHz
6. Output sampling rate :10 Msamples/sec
7.DAC resolution 16 bits
8. Input Voltage range ;1 or 10 Vpp
9.Control matrix linear gain : -20 to 20
10. offset precision 100 microV
4 | Digital Filter Function Features | Shall have at least the following features
1. Filter presets for low pass, high pass, bandpass: Butterworth, Chebyshev, Elliptic, Bessel, Gaussian
2. Adjustable corner frequency shall cover the range 1 Hz to 6 MHz
3. Adjustable filter order from 2nd to 8th
4.Filter Shapes Low pass/High pass/Bandpass/Bandstop/Custom
5. Pass band ripple to be configurable 0.1 to 10dB
6. Sampling rate up to 15 MHz
5 | Phase Meter function | To measure the phase of two signals with very high precision
No of inputs Two inputs signals
| Frequency range | Frequency Range 1KHz to 200 MHz
| Tracing bandwidth | Tracking Bandwidth up to 10 KHz
| Phase precision | Phase Precision 6 microrad/√Hz
Frequency precision10 microHz/√Hz
| PLL tracking bandwidth | PLL tracking 10 Hz to 10 KHz
Real time analysis of Power spectral densities,Allan deviation etc
6 | Lock-in Amplifier Function Features | Shall have at least the following features
1. Dynamic Reserve ≥ 120 dB
2. Internal or External demodulation modes
3. Demodulation Frequency shall cover the range 1mHz to 200 MHz
4. Selectable between rectangular (X/Y mode) or polar coordinate modes
5.Filter steps 6,12,18,24 dB/oct
6.Phase shift precision: 0.001 Deg
7.Built in PID Controller
8. Output gain range; -80 to +160 dB
9. Acquisition model continuous 1MS/s and Snapshot 500 MS/sec
10.Integrated Oscilloscope and data logger mode
7 | FIR Filter builder function | to design and implement various types of filters in time and frequency domain
|  | 1. Frequency response: low pass, high pass, bandpass,bandstop,FIR filters | 387090
|  | 2.Filter coefficients up to 24 bits | 407090 | 244800
3.Impulse response: rectangular,Sine,equation input,custom etc
4.Filter Shapes Low pass/High pass/Bandpass/Bandstop/Custom
5. Filter coefficients upto 14000
6. Sampling rate up to 15 MHz
7. Facility to visualixe filter transfer function,impulse and step response
8 | Arbitrary Waveform generator | to generate custom waveforms with up to 65536 points
1) Max sampling rate: 1GS/
2) output bandwidth: 300 MHz
3) DAC resolution: 16 bits
4) Supports Waveforms; Sine,Guassian,exponential fall and rise,
sinc,equeation editor,custom from file
5) Facility to configure pulse output with up to 250000 cycles of dead time between pulses
6) Two independent channels
7) Phase synchronisation output between two channels
8) Generate signal upto 32 segments
9 | Laser Lock box | Shall have the following features
1. Shall be able to lock laser frequency using Pound Drever Hall locking technique.
2. Shall be able to perform post demodulation filtering.
3. Shall have configurable independent PID controllers for fast and slow feedbacks
4. Shall be able perform scans to find resonance
5. Shall have tap to lock feature to lock to any zero crossing error signal
6. Shall cover the demodulation range from 1mHz to 200MHz
7. Scan frequeny from 1mHz to 1 MHz
10 | Frequeny Response Analyser | to measure Frequency response of a system both in magnitude and phase
1. Frequency range: 10 mHz to to 120 MHz
2.Averaging time: 1 microsec o 10s
3.Sweep points: up to 512 points
4.Linear or logarithmic sweep output
5.Demodulate up to 15th harmonic
6.Noise floor: -120dB @ 120 MHz
11 | Spectrum Analyser | to analyser signals in frequency spectrum upto 250 MHz
1.Frequency span:100 Hz to 250 MHz
2.resolution bandwidth: 1Hz
3.No of inputs: 2 channels
4.Noise floor: -130dB with 1Vpp input
5.Video Filter bandwidth: 10 Hz to 2.4 MHz
6.Integrated Waveform generator of Sine wave upto 250 MHz
12 | Waveform Generator | enables to generate two independent waveforms
1.Sampling rate: 1 GS/s
2.Resolution: 16 bit
3.Frequency range; DC to 250 MHz
4.Five inbuilt waveforms: Sine,square,ramp,pulse and DC
5.FM,Am and PM modulation with internal waveform
6.Pulse width: 4ns
7.Modulation bandwidth: 62.5 MHz
8.Burst Mode: Start,N Cycle ,Gated
9.Sweep time; 1ms to 1 Ks
10.THD: 0.5% (1.5 MHz,5 harmonics)
Note:
1. Bidder shall submit duly filled compliance sheet along with the offer. Offers without compliance sheet will be rejected.
2. Bidder shall clearly mention Make and Model / Part No of the item quoted in the compliance sheet.
3. Bidder shall attach datasheet / product literature supporting the values / remarks in the compliance sheet. Offers without datasheet / product literature will be rejected.
4. All necessary cables / Power adapters usable with standard indian electrical outlet is also part of scope of supply.