## Worksheet
Equipment Name* | Approved Abbreviation | Location | Make (Manufacturer)* | Model* | Serial No* | Funding Agency Type* | Funding Agency Details* | Equipment Description | Department Name* | Facility Name | Basic Cost(Rs)* | Warranty End Date (Valid Till) | Category* | Alert to be sent (Email id)* | PRISM Service Category* | Year of Purchase* | Purpose / Application of Equipment | Operational Status* | AMC End Date (Valid Till) | Access Restrictions | Remarks | Pre-Experiment Support* | Post-Experiment Support* | Charges for Pre Experiment Support | Charges for Post Experiment Support
Aggregate crushing test apparatus | ACT | M 108 | BENT | BENT26 | NA | Internal funding | Institution | Evaluates the strength of coarse aggregates used in road and pavement construction. It measures the material's ability to resist crushing under gradually applied compressive loads, ensuring the aggregates can withstand heavy traffic without degrading.Consists of 150mm diameter steel cylinder,148mm dia. plunger,base plate 200mm square,metal measure- 115mm dia. x 180mm high & tamping rod 16mm dia. x 450mm long | Civil Engineering | TE | 9225 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Buoyancy balance | BB | M 108 | BENT | BENT27 | NA | Internal funding | Institution | To determine the density, specific gravity, and water absorption of solid materials by comparing their weight in air to their apparent weight submerged in water.Electronic balance with 15 kg capacity | Civil Engineering | TE | 46400 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Aggregate impact testing machine | AIT | M 108 | AIMIL | AIM-456 | 162613 | Internal funding | Institution | Measures the toughness of coarse aggregates their ability to resist fracture and breaking under sudden, dynamic loads (such as the pounding impact of heavy vehicles on highways). It ensures structural materials are strong enough for continuous traffic vibrations.Machine is furnished with a quick release trigger mechanism to manipulate an impact with a free fall of a 13.75 kg hammer from a height of 380mm. It is supplied complete with one cylindrical cup, one metal measure of 75mm diameter x 50mm depth and a tamping rod. | Civil Engineering | TE | 23255 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Ductility testing machine | DTM | M 108 | BENT | BENT51 | NA | Internal funding | Institution | Measures a material's ability to plastically deform under tensile stress before fracturing. In road construction and materials science, it evaluates the flexibility, adhesive strength, and cracking resistance of asphalt binders (bitumen) to ensure pavements can withstand traffic and temperature changes.Operation on 220 volts,single phase, A.C. supply. The mould when properly assembled form a briquette specimen of following dimensions. Total length 75.0 ±0.5 mm Distance between clips 30.0 ± 0.3 mm Width at mount of slip 20.0 ± 0.2mm Width at minimum cross-section (half way between clips) 10.0 ± 0.1 mm Thickness throughout 10.0 ± 0.1 mm | Civil Engineering | TE | 62505 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
CBR apparatus | CBR | M 108 | BENT | BENT11 | NA | Internal funding | Institution | Test is a penetration test used to evaluate the mechanical strength and load-bearing capacity of subgrade soils and base course materials in road construction. It directly dictates the required thickness of pavement layers to ensure they can safely support heavy traffic. Load frame, 50 kN capacity with 3 rates of strain i.e. 1.25mm,1.5mm,2.5mm/minute | Civil Engineering | TE | 93600 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Bitumen centrifuge extractor | BCE | M 108 | BENT | BENT59 | NA | Internal funding | Institution | Test determines the exact percentage of asphalt binder (bitumen) in a paving mixture. This test verifies that the asphalt mix meets precise design specifications for road construction, preventing premature cracking, rutting, or bleeding.Electrically operated with variable speeds upto 3600 RPM, bowl capacity- 1500 gm | Civil Engineering | TE | 28080 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Marshall stability apparatus | MSA | M 108 | BENT | BENT53 | NA | Internal funding | Institution | Marshall Stability Test evaluates the resistance of asphalt paving mixtures to plastic deformation, rutting, and shearing stresses. It is widely used in road construction to determine the Optimum Bitumen Content (OBC) for a given aggregate mix and to ensure the pavement can withstand heavy traffic loads.Proving ring capacity - 25 kN. Specimen mould assembly comprising of mould cylinders 10.16cm diameter by 6.35cm height, base plate and extension collars. | Civil Engineering | TE | 82305 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Pensky martens apparatus | PMA | M 108 | AIMIL | AIM-509 | 15621 | Internal funding | Institution | The Pensky-Martens test is primarily used to determine the flash point and sometimes the fire point of petroleum products, lubricants, biodiesel, and other potentially flammable liquids. It utilizes a closed-cup tester to establish the lowest temperature at which a liquid's vapors will ignite when exposed to an ignition source.closed cup is made of brass, the inside of the cup may be turned to a slightly larger diameter above the filling mark and the outside may be tapered above the flange. The flange is about 12mm in width and approximately 3mm in thickness.Lid includes a stirring device, cover, shutter and flame exposure device. The stirring device consists of a vertical steel shaft of 2.5 mm to 3mm diameter and mounted the center of the cup. | Civil Engineering | TE | 21760.73 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Standard tar viscometer | STV | M 108 | AIMIL | AIM-531 | 16683 | Internal funding | Institution | The Standard Tar Viscometer test is used to determine the viscosity and flow characteristics of cut-back bitumen and road oil. It measures the time taken (efflux time) for 50 ml of a heated bituminous sample to flow through a standardized orifice, ensuring the material is fluid enough for construction applications.Operation on 230 volts, 50 Hz, single phase, AC supply. 10mm orifice | Civil Engineering | TE | 17404 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Ring and ball apparatus | RBA | M 108 | BENT | BENT34 | NA | Internal funding | Institution | To determine the softening point of bituminous materials, resins, and waxes. It identifies the temperature at which a material softens enough to allow a standard steel ball to fall through a steel ring filled with the sample.Steel ball having 9.5mm diameter and weighing 3.5 ± 0.05 g. Brass rings-two numbers each having depth of 6.4 mm. The inside diameter at bottom and top is 15.9mm and 17.5 mm respectively. The distance between the bottom of the rings and the top surface of the bottom plate of the support is 25 mm. | Civil Engineering | TE | 9945 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Length gauge | LG | M 108 | AIMIL | AIM-451 | NA | Internal funding | Institution | Used for determining the elongation index test of aggregate.It is the percentage by weight of particles whose greatest dimensions is greater than 1.8 times their mean dimensions. | Civil Engineering | TE | 4608.62 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Thickness gauge | TG | M 108 | AIMIL | AIM-450 | NA | Internal funding | Institution | Used to determine if the aggregate particles are to be considered as flaky, i.e. their thickness is less than 0.6 of their nominal size. | Civil Engineering | TE | 4608.62 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Los-Angeles Abrasion Testing machine | LAAM | M 008 | BENT | BENT49 | NA | Internal funding | Institution | To evaluate the toughness, hardness, and wear resistance of coarse aggregates. It simulates the grinding and impact forces aggregates experience from vehicle traffic and environmental conditions, determining their suitability for road and concrete construction.The drum is rotated at a speed of 30-33 RPM by an electric motor , operation on 440 volts,50Hz, 3 phase, A.C supply | Civil Engineering | TE | 97515 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Hot air oven | HAO | M 108 | KEMI | KOMA 544D | 15854 | Internal funding | Institution | Equipment primarily used to dry aggregates, condition bituminous mixes, and simulate the aging of asphalt binders by providing precise, controlled dry heat.Capacity- 250°C, 220/230 Volts AC | Civil Engineering | TE | 40845 |  | Characterization | renjiths@mgits.ac.in | A | 2016 |  | Working |  |  |  | No | No
Standard penetrometer | SPM | M 108 | BENT | BENT33A | Na | Internal funding | Institution | Used to measure the consistency and hardness of bituminous materials (bitumen and tar). It helps determine the specific paving grade of bitumen to ensure it withstands local traffic loads and temperature variations.Flat bottomed cylindrical metallic dish 55 mm in diameter and 35 mm in depth. If the penetration is of the order of 225 or more, dish of 70mm diameter and 45mm depth is essential.The total moving load is 100 ± 0.25 gm, including the weight of the needle, carrier and super-imposed weights. | Civil Engineering | TE | 7480 |  | Characterization | renjiths@mgits.ac.in | A | 2017 |  | Working |  |  |  | No | No
Modified roughness indicating machine, Merlin | MER | M 108 | AIMIL | AIM 570 | 17170 | Internal funding | Institution | Used to measure road surface roughness,evaluate pavement serviceability and calculate the International Roughness Index (IRI).Device with a probe and two feet which rests on the surface with course of wheel track. T his device is at 1.8 meters seperated and the probe lies at the mid path between two feet. | Civil Engineering | TE | 29818.6 |  | Characterization | renjiths@mgits.ac.in | A | 2018 |  | Working |  |  |  | No | No
Dynamic cone penetrometer | DCP | M 108 | AIMIL | AIM-567 | 17421 | Internal funding | Institution | The primary objective of the DCP test is to calculate the in-situ CBR value of soil without having to extract undisturbed samples for lengthy laboratory testing. To rapidly measure the shear strength and in-place bearing capacity of unbound soils, subgrades, and base courses.The design of the AIMIL DCP is similar to that described by Kleyn, Maree and Savage (1982) in that it uses an 8kg weight dropping through a height of 575mm and a 60° cone having a diameter of 20mm. | Civil Engineering | TE | 20886 |  | Characterization | renjiths@mgits.ac.in | A | 2018 |  | Working |  |  |  | No | No
Indirect tensile strength testing apparatus | IDT | M 108 | Krishna Metal Products- Palarivattom, Kochi | NA | NA | Internal funding | Institution | The Indirect Tensile Strength (ITS) test on Marshall specimens measures an asphalt mixture's resistance to tensile stress. Using the Marshall loading frame, it evaluates pavement durability, fatigue cracking potential, and moisture susceptibility, guiding safe and long-lasting road designs.The sample is placed horizontally between two curved loading strips. A vertical compressive load is applied at a constant deformation rate of 50 mm/min until the sample splits in half along its vertical diameter. | Civil Engineering | TE | 4720 |  | Characterization | renjiths@mgits.ac.in | A | 2020 |  | Working |  |  |  | No | No

## Worksheet 1 (hidden)
NEHU Internal Funds
TCS Funding
SERB, DST, Govt. of India
Central Power Research Institute
MeitY, Govt. of India
BRNS, Govt. of India
VGST, Govt. of Karnataka
AICTE, Govt. of India
Instituion Grant, IIT Madras
ERC Project
NFSG Project
ARDB, Govt. of India
NIT Rourkela
Inode FA Internal
Internal Resource Generation
SRIC Development Fund
Dept of Scientific and Industrial Research DSIR
Ministry of Earth Sciences
NATIONAL ALUMINIUM COMPANY LTD
IISER PUNE
Vinayaka Missions Research Foundation
NITTTR Bhopal
ICAR-CMFRI
Ministry of Textiles
Biotechnology Industry Research Assistance Council
NCBS - TIFR Core Grant
inStem Core Grant
NPMASS, ADA, Govt. of Inda
DRDO, Govt of India
ISRO, Govt. of India
Institution
MoEFCC, Govt. of India
MoFPI, Govt. of India
DST and Institution
Ministry of Education
DAE, Govt. of India
DBT, Govt. of India
Ministry of Petroleum & Natural Gas, Govt of India
Institution Funding (Self Supported)
Office of the Principal Scientific Adviser
Ministry of Defence
Ministry of Road Transport and Highways
Ministry of Health and Family Welfare
Ministry of MSME, Govt. of India
GIA
NCCS Instramural core fund
National Institute of Immunology
NCCS Fund
DOT, Govt. of India
CDOT, New Delhi
National Security Council Secretariat
NDMA Govt of India
DME Government of Karnataka
RUSA, Govt. of India
Government of Kerala
KSCSTE Government of Kerala
Minister of Labour and Employment
Central Water Commission CWC
National Mineral Exploration and Development Trust
Anusandhan National Research Foundation ANRF DST
Dept of Pharmaceuticals NIPER
Ministry of Higher Education
Ministry of New and Renewable Energy
Dept of Atomic Energy
Dept of Health Research Indian council of Medical
Dept of Chemicals and Petrochemicals
Dept of Agriculutre and Farmers Welfare
Dept of DARE DG ICAR
CSIR, Govt. of India
DST, Govt. of India
ICFRE