## Sheet2
Equipment Name | Approved Abbreviation | Location | Make (Manufacturer) | Model | Serial No | Funding Agency Type | Funding Agency Details (Use Comma if multiple) | Equipment Description | Department/Lab Name | Project Title | Rate/Hr | Refernce Website | Category | Equipment Status
BET Surface Area Analyser (BETS) | BETS | AMP | Quantachrome,USA | NOVA 4000e |  | Government | CSIR | Brunauer-Emmett-Teller (BET) surface area and pore size analysis based on the principle of N₂ adsorption-desorption at liquid nitrogen temperature. Present equipment offers determination of BET surface area, mesopore size distribution, pore volume and average pore diameter. | AMP |  |  |  | Research Equipment | Up
Electron beam evaporation unit (Ebeam) | Ebeam | AMP | SN 2237 | HHV |  | Government | CSIR | Electron beam evaporation is used for coating single layer and multilayers of metals, alloys, ceramics. It is an atomic deposition PVD process, done by thermal evaporation where the coating material is melted with the help of electron beam gun in a vacuum of 1*10-⁷ mbar. It has 2 electron guns have 4 crucibles in each. The substrate(on which coating is required) also has heating facility up to 700˚C. | AMP |  |  |  | Fabrication | Up
Fourier Transform Infrared Spectrometer (FTIR) | FTIR | AMP | Thermo | NICOLET 5700 |  | Government | CSIR | This instrument is useful in detecting the chemical bond vibration of organic and inorganic materials.The spectra obtained helps in identifying the chemical nature of the bond and the functional group present qualitatively. It is generally operated with KBr palate for powder sample. With available smart accessories we can record the reflectance spectra of solid powder, gel and liquid materials. | AMP |  |  |  | Characterization | Up
High Temperature Nanoindnetor (HTNano) | HTNano | AMP | MTS | Agilent |  | Government | CSIR | This XP naniondenter works in continuous stiffness measurement (CSM) and control load mode. Berkovich indentor is mainly used for determination of mechanical properties and elastic plastic behaviour of materials. It has nano vision and DCM attachement also for related studies. | AMP |  |  |  | Research Equipment | Up
Magnetic Barkhausen Emission Analyzer (MBEA) | MBEA | AMP | American Stresstech | Rollscan-300 |  | Government | CSIR | The Magnetic Barkhausen emission (MBE) is measured by the instrument Rollscan 300 which is a digital Barkhausen Noise analyzer. This instrument can be used for measurement of Magnetic Barkhausen Emission (MBE) signal. The effect of residual stress, creep and fatigue on MBE signal. The effect of residual stress, creep and fatigue on MBE signal of ferromagnetic materials can be studied. | AMP |  |  |  | Research Equipment | Up
Magnetron Sputtering (MagS) | MagS | AMP | SN 2178 | HHV |  | Government | CSIR | The magnetron sputtering can be used for coating single layer and multilayers of metals, alloys, ceramics and some high temperature polymers such as teflon etc. It is an atomic deposition PVD process done by plasma generation. The unit has the capability of 1*10-⁶ mbar. It has 3 magnetron source with substrate holder having substrate heating and biasing facility up to 700°C and 200V respectively. | AMP |  |  |  | Research Equipment | Up
Melt Spinning (MeltS) | MeltS | AMP | Designed by NML:Fabricated at Vacuum Techniques,Bangalore |  |  | Government | CSIR | The Melt spinning unit involves rapid solidification processing route through melt spinning technique for the preparation of the ribbons directly from the melt. In this technique the molten metal is induction melted and ejected into a rotating copper wheel. The process can be carried out in air or in inert atmosphere. The instrument has the capability of melting 1 Kg of Fe-based alloy. The process parameters include nozzle slit geometry, separation of nozzle slit and copper wheel, wheel speed, ejection melt temperature and pressure. | AMP |  |  |  | Research Equipment | Up
Microhardness (Mhard) | Mhard | AMP |  | Leica VMHT Auto |  | Government | CSIR | Microhardness unit is equipped with Vickers and Knoop indentor. The load capacity is from 1 gm to 2000 gm in discrete increments. It is used for measurement of static hardness of different phases in the materials and coatings. | AMP |  |  |  | Research Equipment | Up
PECVD unit | PECVD | AMP |  |  |  | Government | CSIR |  | AMP |  |  |  | Research Equipment | Up
Phased Array/TOFD | TOFD | AMP | RD Tech, Olympus | OMNISCAN |  | Government | CSIR | Ultrasonic system with Time-of-flight diffraction (TOFD) is a technique that uses two probes in pitch-catch mode and detect and records signals diffraction from defect tips allowing both detection and sizing. The TOFD data is displayed in a gray scale B-scan view. TOFD offers wide coverage and amplitude-independent sizing complying with the ASME 2235 code.One-line scan for full- volume inspection.Setup independent of weld configuration. Very sensitive to all kinds of defects and unaffected by defect orientation. | AMP |  |  |  | Research Equipment | Up
Plasma Spray Coating System | PPIN | AMP | MILMAN, PUNE | PLASMON 600 |  | Government | CSIR | Pulsed plasma ion nitriding(PPIN)is used to modify the surface of steel components through controlled nitrogen diffusion process to achieve high hardness, corrosion, fatigue,and wear resistance. PPIN is a modern and environmentally clean method of incorporating nitrogen in steel components. This makes it possible to nitride the steel surfaces even white layer thus, reducing finishing cost of the component. By controlling the pulse duration and repitition time, arc discharge can be effectively suppressed, treatment temperature can be varied without altering plasma parameters and phase contitution of the layer can be altered. Very complex shapes can be treated successfully and the actives species can also enter blind holes, producing completely uniform modified layers on all parts in a heavily loaded chamber. | AMP |  |  |  | Research Equipment | Up
Ultrasonic C Scan Imaging (UTCS) | UTCS | AMP | Physical Acoustics |  |  | Government | CSIR | Ultrasonic c-scan system is used to identify the cleanliness of materials in terms of inclusions, defects etc through 2D and 3D images. The system is used to scan a material in a noncontact way by immersing it in a tank filled with water. Frequency range of the system is upto 20MHz. This system can be used to find the macrostructure of cast materials, inclusion distribution at different depths, pit size and depth distribution, central quality evaluation of cast billets & blooms etc. | AMP |  |  |  | Research Equipment | Up
Fluidised Bed Jet Mill | FDJM | MER | Montech Instruments, Chandigarh |  |  | Government | CSIR | The high temperature fluidized bed reactor is a sophisticated pilot unit for experimental studies of reactions involving reduction of ores to metallic state using gases like CO and LPG and neutral gases like N₂ and argon. The gas flow may be metered and manually controlled with the state of the art mass flow meters giving direct digital output of mass flow rate, pressure and temperature. Over an experimental run, gases used may be preheated before entry to the reactor. This hot gas mixture is further heated to the process temperature by passage through a bed filled with alumina balls in the INNER TUBE that has the Zirconia distribution plate at the top. The ore bed is on the distributor plate and gases flowing through its fluidize it. THe inner tube with distributor plate is inside the main reactor tube. The fluidized ore column is roughly in the middle of the hot reaction zone of the three zone main furnace. The zones can be independently controlled/trimmed to achieve a suitable temperature profile in the main reactor. The reaction mixture can be sampled by a unique device inserted from the top flange of the reactor into the reactor into the reaction zone. Full protection against over pressure in gas lines is achieved by using special msafety valves that have adjustable trip points at present to set 20 psi. Non return values are used in each gas line to prevent reverse flow. The reacted contents can be downloaded into the lower chamber by sliding open the special plate valve at the bottom. This chamber is double walled water cooled housing with two thermocouples for temperature monitoring and a viewing window for observation. This window has tempered glass. | MER |  |  |  | Research Equipment | Up
High Temperature Muffle Type Furnace | HTFM | MER |  |  |  | Government | CSIR |  | MER |  |  |  | Research Equipment | Up
Inductive Coupled Plasma Optical Emission Spectromter (ICPOES) | ICP-OES | MER | Vista-MPX | Varian |  | Government | CSIR | Inductively coupled plasma optical emission spectrometry (ICP-OES), is an analytical technique used for the detection of trace metals(ppb). It is a type of emission spectroscopy that uses the inductively coupled plasma to produce excited atoms and ions that emit electromagnetic radiation at wavelengths characteristics of a particular element. The intensity of this emission is indicative of the concentration of the element within the sample. Unlike atomic absorption spectroscopy, which can only measure a single element at a time ICP-OES has the capability to scan for all elements simultaneously. This allow rapid sample processing. | MER |  |  |  | Research Equipment | Up
Muffle type Reduction Pusher Furnace | Muffle | MER |  |  |  | Government | CSIR |  | MER |  |  |  | Research Equipment | Up
Optical Emission Spectrometer (OES) | OES | MER | Vista-MPX | Varian |  | Government | CSIR |  | MER |  |  |  | Research Equipment | Up
Particle Size Analyser (PSA) | PSA | MNP | MALBERN | Master sizer 2000 |  | Government | CSIR | Laser diffraction Particle size Analyser Consists of wet dispersion system Hydo2000MU built-in recirculation pump, ultrasonic probe, mechanical stirrer, dry dispersion system Sirocco2000 and an application software. It works with the optical principle of laser diffraction and laser scattering, and provides the particle size distribution in volume percent. THe d₁₀,d₅₀ and d₉₀ indicating the size particles present below 10%,50% and 90% respectively in an analysis is computed by the software. Software provision exists to exhibit particle size distribution against a size scale in ASTM mesh size, Phi scale and micron. | MER |  |  |  | Research Equipment | Up
Vacuum Arc Melting (VAM) | VAM | MER |  |  |  | Government | CSIR |  | MER |  |  |  | Research Equipment | Up
Zeta Probe Analyser (ZPA) | ZPA | MER | Colloidal Dynamics USA |  |  | Government | CSIR | Zetaprobe(Make: Colloidal Dynamics USA) is used for measuring zeta potential of particles which is an extremely important parameter across a wide range of industries including brewing, ceramics,pharmaceuticals,medicines,mineral processing and water treatment. It gives a measure of the stability of the colloidal system. Its measurement brings detailed insight into the dispersion mechanism and is the key to electrostatics dispersion control. This equipment measures the zeta potential in colloidal suspensions and emulsions. The equipment works on multi-frequency electro acoustic method. Its automatic titration facility facilitates determination of isoelectric point. It has provision for correction for particle size and double layer distortion. | MER |  |  |  | Research Equipment | Up
Classification & Gravity Separation Facility (GSF), Bench Scale | GSF | MNP |  |  |  | Government | CSIR |  | MNP |  |  |  | Research Equipment | Up
Crushing and Grinding Facility (CGF) | CGF | MNP |  |  |  | Government | CSIR |  | MNP |  |  |  | Research Equipment | Up
Instrumental Particle Size Analysis Facility (IPSA) | IPSA | MNP |  |  |  | Government | CSIR |  | MNP |  |  |  | Research Equipment | Up
Mercury Porosimeter PA140/440-MNP | MERP | MNP | M/s PMI, USA | PMI 60K |  | Government | CSIR | Mercury Porosimeter is used to evaluate the porosities in samples, which can be metals and ceramics or composites in the form of either solid compact(sintered/unsintered)or in powder form. | MNP |  |  |  | Research Equipment | Up
Shatter Index (SHI) | SHI | MNP |  |  |  | Government | CSIR |  | MNP |  |  |  | Research Equipment | Up
Surface Tensiometer | SURT | MNP | Germany | M/s Kruss K100 |  | Government | CSIR | The tensiometer is used to measure the contact angle of solids, powders and fibre bundles. Surface free energy calculation can be determined from the contact angle data. Surface tension of liquids, interfacial tension between two liquids can be measured. The measurement method involves both ring and plate (with highly accurate contours) mode operation. The measurements can be made between temperatures varying from RT to 130° C. | MNP |  |  |  | Research Equipment | Up
Tumbler Index Machine-MNP |  | MNP |  |  |  | Government | CSIR |  | MNP |  |  |  | Research Equipment | Up
Electron Probe Micro Analyser (EPMA) | EPMA | MTE | JEOL,Japan | JXA-8230 |  | Government | CSIR | Electron Probe Micro Analyser (EPMA)is high resolution elemental analysis instrument that combines wavelength dispersive X-ray spectrometers(WDS) and energy dispersive (EDS) spectroscopy. It can analyse elements both qualitatively and quantitatively. Depending on requirement W and LaB6 filament can be used. Point, line and area analyses are possible, also elemental distribution can be obtained by map analysis. THe minimum probe size is 0.5 micron (region of analysis). | MTE |  |  |  | Research Equipment | Up
High Strain Rate (HSR) Mechanical Testing System | HSR | MTE | Instron | VHS |  | Government | CSIR | Recent automobile saftey regulations and subsequent development of analytical models for automobile crash events have mandated the need for realistic constitutive data at high strain rates. This data is required as inputs for computational design models. If accurate numerical simulations can be obtained from computational models and high strain rate data, the expensive crash tests may be dramatically reduced. This not only translates into significant cost savings, but also will allow the automobile manufacturer to make minor adjustments in their designs in a more timely and efficient manner. Keeping in view, a state of the art high strain rate test facility is made available for the industrial application. | MTE |  |  |  | Research Equipment | Up
High Temperature Softening Melting (HTSM) | HTSM | MTE | Indigenous Con Com Consultants, Ranchi |  |  | Government | CSIR | The softening-meltdown properties (SM)of ferruginous material have important influence on the furnance operation including pressure drop, distribution of hot ascending gas,fuel rate, hot metal quality, etc. Accordingly,it is desirable to assess the suitability of the iron bearing raw material prior to their use in the blast furnace. The HTSM equipment tests the ferruginous raw material in the laboratory under blast furnace simulated conditions and parameters those can be calculated/inferred by the test are softening temperature(Ts), melt down temperature(Tm) , resistance to gas flow, bed shrinkage and % Non-dripped material(NDM) | MTE |  |  |  | Research Equipment | Up
Hot Forging | FRG | MTE |  |  |  | Government | CSIR |  | MTE |  |  |  | Research Equipment | Up
Hot Rolling with Salt Bath Treatment | HRM | MTE |  |  |  | Government | CSIR |  | MTE |  |  |  | Research Equipment | Up
Instron Tension Torsion Test (TTT) Machine | TTT | MTE |  | Instron 8850 |  | Government | CSIR | This unique facility is the first of its kind in the country. Deformation behaviour of materials can be studied under combined axial-torsional load/strain paths with or without applying an internal pressure on tubular specimens. The two load paths can be controlled either in-phase or out of phase. The information generated is useful for understanding the deformation behaviour of materials under multiaxial loading conditions in stress state situations analogues to those of actual components. THe system has an axial capacity of +/-100kN and torsional capacity of 1000Nm. THe maximum internal pressure that can be applied using a special unit is 600 bar. THe system also has bi-axial extensometer for precise control/measurement of axial and shear strains. | MTE |  |  |  | Research Equipment | Up
Nanotribometer (NTB) | NTB | MTE | CERT(USA) | UMT-2 |  | Government | CSIR | Nanotribometer is one of the most important tools for evaluation of nano-tribological properties (nano-wear and nano-scratch) of bulk materials, thin films, melt spun ribbions and conventional materials of metallic and non-metallic compositions. The system is capable of measuring nanotribological parameters from linear reciprocating tests and rotational tests. Test should be possible with pins and balls of miniature dimensions, with loads and displacements measured in the relevant range with nano-to micro- resolution. For linear reciprocating module stroke length is upto 75 mm and for rotation module the speed varies from 10 to 1000rpm, clockwise, counter clockwise, cyclic and gyratic motions. The test types are Ball-on-test specimen or Pin-on-test specimen. The force measurement range is 50m to 0.5 N with force measurement resolution is 10mN. The wear deapth can be measured upto 200m measurable wear with linear resolution 25 nm and for 2mm or more measurable wear the linear resolution is 0.5mm. High magnification optical microscope is attached to view the sample and wear track without removing sample. The system is also attached with profilometer(DEKTAK-XT)for wear profile measurment. | MTE |  |  |  | Research Equipment | Up
Pendulum Impact Tester RKP 450 | CVN | MTE |  | Zwick-Roell |  | Government | CSIR | Conventional impact test machines are used for estimating the notch toughness of materials, which in turn is used as the material screening criteria in manufacturing industries. The instrumented impact test machines are a special variety,in which, apart from the conventional notch toughness, the system would capture the load and displacement data within few milliseconds of the hammer contact with the specimen. This data can be used to determine the dynamic fracture resistance of the material such KID and the ductile to britlle transition temperature (DBTT) of metals and alloys. The system is equipped with a temperature bath in the range of +150° C to -80°C. Dynamic fracture characterization of materials in awide spectrum of temperature is therefore possible. The system is calibrated with NIST (National Institute of Standards and Testing-USA) standard reference specimens. | MTE |  |  |  | Research Equipment | Up
Pi On Disc (POD) Wear Tester | POD | MTE | DUCOM,Bangaluru | CM-9065 |  | Government | CSIR | Pin-on-Disc Wear Tester is used to study the material loss due to frictional forces between the two moving bodies in contact. The system is capable of simultaneously measuring the friction coefficient and wear loss with time or sliding distance under varying load. The lubricability of the fluid can be evaluated. | MTE |  |  |  | Research Equipment | Up
SEM with EDS and EBSD | SEM-EDS | MTE |  | FEI Nova nanoSEM-430 |  | Government | CSIR | FEI Nova NanoSEM 430, an ultra-high resolution field emission SEM which is specifically configured to get the most information out of the largest selection of samples, down to the nanometer level. The Nova NanoSEM 430 has the capabilities of very low kV characterization as well as its analytical capabilities. Ultra-high resolution results can be obtained on a variety of challenging materials such as metals, magnetic materials, nano materials. Available nanoanalysis capabilities such as EDS and EBSD with Nova NanoSEM 430 wil provide a powerful solution for the most demanding characterization needs at the nanoscale. | MTE |  |  |  | Characterization | Up
SEM with FEG-EDS | SEM-FEG | MTE |  |  |  | Government | CSIR |  | MTE |  |  |  | Characterization | Up
TEM -FEG | TEM-FEG | MTE |  | JEOL JEM 2200FS |  | Government | CSIR | The JEOL JEM 2200FS is an advanced and unique transmission electron microscope(TEM) that covers a range of imaging mode and spectroscopic applications. The state of art analytical equipment, eqipped with schottky FEG need no dark room and produce rotation-free images. The main frame of the TEM is mounted on a passive type air mount and supported with oil free vacuum pumping systems. The instruments has in-column W filter that makes it suitable for EELS study. The microscope is also equipped with STEM with minimum probe size 0.2 nm. STEM mode benifits with HAADF, with 80mm2 SDD detector ideal for point analysis and elemental mapping analysis.TEM images are recorded didgitally on a Gatan 2k *2k CCD camera. The TEM has other facilities like HR-TEM, NBD and CBED. | MTE |  |  |  | Characterization | Up
Thermo Mechanical Simulation | GLEEBLE | MTE | Dynamics Systems Inc.(USA) | GLEEBLE 3800 |  | Government | CSIR | The thermo-mechanical/physical simulation is one of the most important aspects in the development of materials processing methodology. This involves exact reproduaction of the thermal and mechanical processes employed in actual component fabrication. The GLEEBLE thermo-mechanical simulation system is a dynamic testing machine that can be used for the simulation of wide variety of metallurgical processes such as single and multi-pass rolling,extrusion, forging, continuous casting, hot tensile and compressive testing, HAZ in welding, thin strip annealing etc. In addition, the system can be effectively used for hot ductility testing, hot/warm compressive and tenslie testing, SICO test, THermal cycling/heat treatment simulations and dilatometry/phase transformation(CCT)studies. The system uses direct resistance heating of the samples and,therefore, avoids the use of a furnace and inaccuracies involved thereof. A closed-loop hydraulic servo system provides precise control of all mechanical variables. The system can operate in the range from room tempertaure to 1700°C. The thermocouple data acquisition rate can be varied, with a resolution of 1°C , depending upon the requirements. Typically, a sample of materials is heated and mechanically worked while various performance parameters of interest are measured and recorded for later analysis. After the simulation or test is done, the microstructure of the material may be examined and analyzed to see the effects of the process schedule adopted. | MTE |  |  |  | Research Equipment | Up