## Sheet2
Equipment Name | Approved Abbreviation | Location | Make (Manufacturer) | Model | Serial No | Funding Agency Type | Funding Agency Details (Use Comma if multiple) | Project Title | Equipment Description | Department Name | Facility Name | Basic Cost(Rs) | Refernce Website | Equipment Status | Category | FOV Cost(Rs) | CIF Cost(Rs) | Any Other Cost(Rs) | Supplier | Supplier Email Id | Service Provider | Service Provider Email Id | Alert to be sent ( Email id )
Field Emission Scanning Electron Microscope | FESEM | Room no.G12, CMCT extension; ground floor | Carl Zeiss | Gemini 500 | 8203017091 | ARCI operational fund |  |  | This instrument is used for studying sample surfaces at high magnifications due to a hot Schottky field emission (FE) gun. | Centre for Materials Characterization and Testing | Field emission scanning electron microscope |  | www.arci.res.in | UP | [8 empty] | ramakrishna@arci.res.in
High flux X-ray diffraction unit | HF-XRD | Room no.G12, ground floor, CMCT | RIGAKU Corp., Japan | SmartLab 2018 | BD68000047-01 | ARCI operational fund |  |  | High Flux X-ray diffraction (XRD) unit is equipped with 9kW rotating Cu anode which provides nearly one order higher flux on the sample than the sealed tube X-ray source. This XRD machine can be used to carry out angular dispersive scans with 2 from 5 to 160 deg. and an angular resolution of 0.01 degrees. Specially designed sample holders are available for mounting of solid samples and powder samples. With aid of auto sampler robot, a maximum of 48 powder and 24 bulk samples can be measured with operator intervention. | Centre for Materials Characterization and Testing | Hiflux x-ray diffraction unit |  | www.arci.res.in | UP | [8 empty] | sureshkoppoju@arci.res.in
High intensity Micro area XRD unit | None | Room no.G104, ground floor; CMCT | RIGAKU Corp., Japan | RAPID-II-D/MAX | BD63000450-01 | ARCI operational fund |  |  | The microficus XRD is equipped with high intensity microfocus rotating anode x-ray source. The system also has a highly sensitive and large image plate based 2-dimensional (2D) detector, which ensures a full scan from -47 to +163 degree for 2theta in a single exposure of very small duration. It has both Cu and Cr targets the with provision to bring the beam spot size down to 10 micron. The system can be operated in reflection, transmission and glancing incidence configuration. It can perform wide range of studies like phase analysis, texture, residual stress, trace phase detection (for phase fraction of 0.1% or less) in micro/macro-area apart from thin film analysis in the glancing incidence mode. | Centre for Materials Characterization and Testing | Microfocus XRD unit |  | www.arci.res.in | UP | [8 empty] | joydip@arci.res.in
Universal testing machine | UTM | Room no. G107, ground floor, CMCT | Instron | 5584 | 56732 | ARCI operational fund |  |  | This unit is used to study the mechanical properties of materials under tension and compression. The sample to be tested is gripped between the lower fixture of the unit and the upper crosshead and the desired load is applied. The computer interface measures the crosshead movement and the extension/compression of the sample. The unit is equipped with a video extensometer so that the change in dimension of the sample can be measured without any sensors coming in contact with the sample. Extra ports are available for additional devices such as an ultrasonic detector to study grain boundary migration. | Centre for Materials Characterization and Testing | Universal Testing Machine |  | www.arci.res.in | UP | [8 empty] | nravi@arci.res.in
Non-contact optical profiler | None | Room no. G19, ground floor, CMCT | Zygo Corporation | New View 6200 | 06-22-5324 | ARCI operational fund |  |  | The system is used for estimating surface roughness and other surface parameters, as also to create 3D profiles of surfaces. It uses white light interferometry to map specimen surfaces. A beam splitter is used to split the incoming beam into two. One part goes to a reference mirror that is optically smooth while the other part scans the sample surface. The interference pattern between the two beams (that is, the one coming from the reference mirror and the other from the sample) is used to recreate the surface profile. The resolution of the unit is 0.1 nm and hence optically smooth surfaces can be studied. | Centre for Materials Characterization and Testing | Non-contact optical profiler |  | www.arci.res.in | UP | [8 empty] | sureshkoppoju@arci.res.in
Small angle X-ray Scattering unit | SAXS | Room no. G16; CMCT extension; ground floor | Xenocs | Genix | G09-12-33/34 | ARCI operational fund |  |  | The laboratory small-angle X-ray scattering (SAXS) unit is used to determine the size and shape of precipitates in a matrix. The SAXS system at CMCT in ARCI is equipped with dual energy (Mo and Cr) sources and has been designed with the aim of studying heterogeneity in high-Z metallic structural materials such as steels (Fe alloy). With the combination of Mo and Cr energies, various camera lengths (up to a maximum of 2400 mm) and an area detector, three decades of q range can be achieved, from 0.024 to 14 nm-1. In real space, the probing periodic distance is a maximum of 261 nm. Precipitates with sizes in the range from sub-nanometre to about 150 nm can be studied using the system. A flexible sample mounting stage permits measurements on a wide range of materials in transmission geometry. | Centre for Materials Characterization and Testing | Small angle X-ray scattering unit |  | www.arci.res.in | DOWN | [8 empty] | sureshkoppoju@arci.res.in